|  | Medical Devices Control Division Ministry of Health Malaysia Level 5, Plot 3C4, No. 26, Jalan Persiaran Perdana, Precint 3, 62675 Putrajaya Malaysia Tel: 603-8885 0778 / 603-8885 0773 Fax: 603-8885 0758 |
| --- | --- |

Medical Device / Equipment VOLUNTARY RECALL

Date Issued : 20 AUGUST 2010                         Ref:MDB/R/2010/009

| IMMEDIATE ACTION | √ |
| --- | --- |
| ACTION |  |
| UPDATE |  |
| INFORMATION REQUEST |  |

*The dissemination of this information is based on the information gathered via their Voluntary Recall Summary and electronic mail from Johnson & Johnson Sdn. Bhd.

| PRODUCT | A limited number of 1 DAY ACUVUE TruEye contact lenses (Narafilcon A)-manufactured in Ireland (Refer to the attachments for details of the affected lot numbers) |
| --- | --- |
| CLASS | N/A |
| USE | Contact Lens |
| SOURCE OF MEDICAL DEVICE RECALL / ALERT | Letter and electronic mail from: Johnson & Johnson Vision Care Malaysia. |
| RECALLING / ALERTING FIRM | Johnson & Johnson Vision Care Unit 11.01, level 11, Uptown 2, 2 Jalan SS21/37, Damansara Uptown, 47000 Petaling Jaya Malaysia. |
| REASON FOR RECALL/ALERT | A limited number of customer complaints associated with affected lots of Narafilcon A Contact lens. Affected lots are those manufactured in Ireland and distributed in Malaysia. Some reported complaints described unusual stinging or pain upon insertion with Narafilcon A contact lenses, and some also associated with ocular redness. Investigations showed that a processing agent was not properly extracted from the lens. |
| SCENARIO IN MALAYSIA | Johnson & Johnson Vision Care conducted a Voluntary Recall on their Narafilcon A contact lenses which bear a product name: 1-Day Acuvue TruEye. All trial lenses with the affected parameters will be recalled from eyecare practitioners and all consumers. Please see attachments for details. |
| ACTION | *any individual possessing the abovementioned device are advised to respond to this communication immediately. |
| RECOMMENDATION | Users of the abovementioned device should contact the distributors/supplier of this device/product and inform the Medical Devices Control Division, Ministry of Health providing the following information:- a. Contact person and contact number b. Numbers of units available c. Name of the supplier(s) |
| CONTACT/ENQUIRIES | Ng Weng Leong Johnson & Johnson Vision Care Unit 11.01, level 11, Uptown 2, 2 Jalan SS21/37, Damansara Uptown, 47000 Petaling Jaya Malaysia. T: 603 7710 0068 ext 217 |